## Slide 1
MoHFW /CDC / BD Partnership Project
LABS FOR LIFE PROJECT - INDIA
Equipment Management

## Slide 2: Complete Blood Count Analyzers
These analyzers are used to count the number of different types of red and white blood cells, blood platelets, hemoglobin, and hematocrit levels in a blood sample.
MoHFW /CDC / BD Partnership Project

## Slide 3: Types of Cell Counters
MoHFW /CDC / BD Partnership Project
3-Part Differential Cell Counters
  Provides neutrophils, lymphocytes and mononuclear cell counts
  Unable to distinguish between monocyte, eosinophils, and basophils.
5-Part Differential Cell Counters
  Able to determine granularity, diameter, and inner complexity of blood cells.
  Able to differentiate all 5 population of WBCs
  Some cell counters are able to detect and abnormal cell populations
Erythrocytes
Lymphocytes
Monocytes
Neutrophils
Basophils
Eosinophils
Platelets

## Slide 4: Principle of operation of Cell counter
Impedance Principle:
Cells are non-conductive
Electrical resistance caused by cells is proportional to their volume /size.
Various blood cells could be differentiated based on this principle
MoHFW /CDC / BD Partnership Project

## Slide 5: Principle of operation of Cell counter
Diluted blood is passed between two electrodes through an aperture
Impedance changes as a cell passes through is monitored.
A lytic reagent is added to the blood solution to selectively lyse the red cells (RBCs) leaving only white cells (WBCs), and platelets intact.
The platelet count is easily separated from the WBC count by the smaller impedance spikes they produce in the detector due to their lower cell volumes
Impedance principle is used for CBCs and three-part WBC differentials
MoHFW /CDC / BD Partnership Project

## Slide 6: A red cell passes through RBC aperture
Sensing Zone
Red Blood Cell
The Impedance principle
Oscilloscope
Oscilloscope

## Slide 7: A red cell passes through RBC aperture
Sensing Zone
Red Blood Cell
Oscilloscope
Oscilloscope
The Impedance principle

## Slide 8: A White cell passes through WBC aperture
Oscilloscope
Sensing Zone
Neutrophil
Oscilloscope
The Impedance principle

## Slide 9: A White cell passes through WBC aperture
Oscilloscope
Sensing Zone
Neutrophil
Oscilloscope
The Impedance principle

## Slide 10: A White cell passes through WBC aperture
Oscilloscope
Sensing Zone
Neutrophil
Oscilloscope
The Impedance principle

## Slide 11: Applying Thresholds to separate Plt from RBC
20 fl
2 fl
Base line
RBC of various
sizes
Platelets

## Slide 12: 3-Part Differential Histogrm
Lymphocyte
Monocyte
Granulocyte
256 channel high resolution WBC histogram
display cell population data between 35 and 450 fl

## Slide 13: Coulter WBC histogram
Lymphocytes
Monocytes
Eosinophils
Neutrophils
Basophils

## Slide 14: Principle of operation of 5-part Differential Counters
Besides using impedance principle, 5-part differential cell counters employ flowcytometric principles
Through hydrodynamic focusing cells are passed through an aperture in single file
Cells are subjected to laser beam which assists in:
  Identifying granularity of the cells
  Identify nuclear structure
Based on size, granularity and nuclear structure, cells are separated in 3-dimentional space
All five types of WBCs are identified based on their typical position in 3-dimentional space
MoHFW /CDC / BD Partnership Project

## Slide 15: Hemoglobin Measurement
MoHFW /CDC / BD Partnership Project
ADC
Signal
Processor
HGB Lamp
WBC Bath
Sample
HGB Sensor &
PRE-AMP
Absorbance= log 10 ( VR / Vs )
Where VR = Reference voltage
  Vs = Sample voltage
Derivation

## Slide 16: Cell counter – Key components
MoHFW /CDC / BD Partnership Project
Single Sample port
Tube rack station
Tube cap piercer
Typical report

## Slide 17: Maintenance of Cell Counters
MoHFW /CDC / BD Partnership Project
Follow your manufacturer’s instructions manual for specific maintenance needs of your analyzer – instructions given here are only indicative
Always use PPE while performing maintenance
Keep the cell counter surface and surroundings clean using 10% bleach.
Most of computerized cell counters have computer-controlled maintenance menu. Users must be familiar with these
Inspect the bottom of washing head for salt build up – should be wiped off with a damp cloth or wiper
Check tubing system and look for any liquid leakage. If you experience leakage, contact authorized technician.
Perform daily start-up and shut-down cycles – these cycles are important for auto-cleaning of tubing and apertures and hence in reducing blank values
Check reagent status and change if required
Empty waste container before starting the day
Waste contains human origin substances representing biohazard.
    Neutralization of biohazard waste:
    Put 2 ml per liter of hypochlorite solution into the waste. Close the cap and
    Shake the container., after 1 hour you can dispose of the Waste liquid into the drain.

## Slide 18: Example of Weekly Maintenance of Cell Counters
MoHFW /CDC / BD Partnership Project
Follow your manufacturer’s instructions manual for specific maintenance needs of your analyzer – instructions given here are only indicative
Cleaning of Probe washing head:
  Salt build-up on the lower surface of wash head may cause malfunction during operation leading to carryover between samples.
  Use a soft cloth or wiper dampened with water to clean this area.
Inspect aspiration / dilution syringes for leaks or salt buildup
Clean sample auto-loader station and track
Probe wash head

## Slide 19: Calibration of cell counters
MoHFW /CDC / BD Partnership Project
Hematology calibrators have short expiry and very short open vial stability.
Ensure traceability of calibrators to ICSH.
Some 3 part equipment calibrators have a defined Uncertainty of Measurement and give a range for the target.
Calibrators with more accurate values give a single, equipment specific target value.
It is important to ensure adequate equipment performance in terms of precision, background values and carry over values before performing CBC analyzer calibrations.
Both closed and open modes should be calibrated whenever 2 modes are available.
Calibration should be followed by verification QC to verify acceptable performance
Automated hematology analyzers shall be calibrated using calibrator provided by manufacturer: Controls shall not be used for calibration.