## Slide 1: It Begins with the Right Chart Labs for Life Project Quality Control Training

## Slide 2: Assigned ≠ Observed
“Theoretically, laboratory QC is easy. Simply calculate the mean and SD for a set of data and use one or more QC rules to warn when data fall outside the acceptable range. Unfortunately, many pitfalls exist that may lead to incorrect conclusion about method performance.”
Brooks, Zoe C. Performance-Driven Quality Control, 2001, p 70.

## Slide 3: 2 reasons for more data points to lie farther from the mean
A change in accuracy or precision
  Systemic changes; imprecision and shifting accuracy
  More than 5% of the measurements will exceed ± 2SD
The mean and/or SD are not assigned correctly on the L-J chart
  Assigned values ≠ Observed values
  More than 5% of the measurements will exceed ± 2SD
Notes: We saw in the previous slide about imprecision or shifting accuracy and more data points falling in the SD areas more than the specified percentages. Such events denote errors in the analytical system. However, there is a more critical and often overlooked reason for his happening. Carelessness in assigning the right values on the chart. Let us understand more

## Slide 4
  Assigned Mean ≠ Observed Mean
  Assigned SD is ≠ Observed SD
Rendering QC rules ineffective.
Notes: 2 cardinal errors made by even well meaning operators!

## Slide 5
MEAN obs
MEAN ASSIGNED
Notes: The mean assigned on the chart is depicted by the black dotted line. 1SD green lines, 2 Sd the red lines. The QCs are being run diligently, but all data points are above the mean. The expected Gaussian is here, but the actual Gaussian is there. So what is the problem? A wrong mean. Conseqently what would the lab do? Look at these points. Multiple points falling outside 2 sd, even 3 sd and continuous false rejections

## Slide 6
False Reject
True Accept
False Accept
93
3
90
3

## Slide 7
Notes: Assigned mean ≠ observed mean

## Slide 8
Mean assigned
Mean observed
False Rejects
False Accepts
It’s a Fact – this is WHERE WE ARE
This is WHERE WE EXPECTED TO BE – based on the assigned values of the chart
Notes: Assigned mean ≠ observed mean

## Slide 9
  Assigned SD ≠ Observed SD
Rendering our QC rules ineffective.

## Slide 10
MEAN obs
SD obs
SD assigned
Notes: False accepts, too wide SDs. Wrong chart

## Slide 11
Notes: SD assigned is greater than observed SD

## Slide 12
5
SD observed
SD assigned
WHERE WE ARE
False Accepts
False Accepts
True Accepts
Notes: SD assigned is greater than observed SD

## Slide 13
Notes: SD assigned is less than observed

## Slide 14
6
SD observed
False Rejects
False Rejects
SD assigned
Notes: SD assigned is less than observed

## Slide 15: For Each QC Data Point We Need to Decide…
SYSTEM IS STABLE
SYSTEM IS UNSTABLE – SIGNIFICANT CHANGE
NO RULE VIOLATION
RULE VIOLATION
OCCURRED ON L-J CHART
True Accept – system is stable and no QC rule violations occurred - report patients
False Accept – system is unstable and no QC rule violations occurred –wrong results may be reported
False Reject– system is stable and a QC rule violation occurred due to inherent randomness- waste of time, effort, and
True Reject – system is unstable and a QC rule violation occurred – begin investigation and troubleshooting

## Slide 16: Quiz Time!!!

## Slide 17: It Begins with the Right Chart
Pitfalls encountered when assigned values that do not reflect current observed values on our charts
Does this look like an error to you?

## Slide 19
Chart: LJ 1. Feb 2016 - AST LI Right chart
 | 70 | 68 | 60 | 59 | 72 | 74 | 70 | 73 | 74 | 72 | 85 | 77 | 69 | 69 | 72 | 70 | 66 | 65 | 80 | 71 | 70.4 | 79 | 79.3467 | 65 | 71 | 68 | 70 | 72 | 70 | 70
Chart: LJ 2. Feb 2016 -AST L1 (Wrong chart)
 | 70 | 68 | 60 | 59 | 72 | 74 | 70 | 73 | 74 | 72 | 85 | 77 | 69 | 69 | 72 | 70 | 66 | 65 | 80 | 71 | 70.4 | 79 | 79.3467 | 65 | 71 | 68 | 70 | 72 | 70 | 70
Notes: Let us see the same values plotted with 2 means. The right mean is 72. if the lab assigns 67 as shown in the second graph, what will happen? Look at the 3rd and 4th values. It is a 2:2s, but will the lab pick it up? No. False acceptance. Look at eh 12th value. It was a 1: 2s warning which became a 1:3s rejection

## Slide 20
Chart: LJ 1. Feb 2016 - AST LI Right chart
 | 70 | 68 | 60 | 59 | 72 | 74 | 70 | 73 | 74 | 72 | 85 | 77 | 69 | 69 | 72 | 70 | 66 | 65 | 80 | 71 | 70.4 | 79 | 79.3467 | 65 | 71 | 68 | 70 | 72 | 70 | 70
Chart: LJ 2. Feb 2016 -AST L1 (Wrong chart)
 | 70 | 68 | 60 | 59 | 72 | 74 | 70 | 73 | 74 | 72 | 85 | 77 | 69 | 69 | 72 | 70 | 66 | 65 | 80 | 71 | 70.4 | 79 | 79.3467 | 65 | 71 | 68 | 70 | 72 | 70 | 70
Notes: Wrong SD. Too wide. Right SD 5. Wrong SD 10. No warnings, no rejections. False accepts.

## Slide 21
Chart: LJ 1. Feb 2016 - AST LI Right chart
 | 70 | 68 | 60 | 59 | 72 | 74 | 70 | 73 | 74 | 72 | 85 | 77 | 69 | 69 | 72 | 70 | 66 | 65 | 80 | 71 | 70.4 | 79 | 79.3467 | 65 | 71 | 68 | 70 | 72 | 70 | 70
Chart: LJ 2. Feb 2016 -AST L1 (Wrong chart)
 | 70 | 68 | 60 | 59 | 72 | 74 | 70 | 73 | 74 | 72 | 85 | 77 | 69 | 69 | 72 | 70 | 66 | 65 | 80 | 71 | 70.4 | 79 | 79.3467 | 65 | 71 | 68 | 70 | 72 | 70 | 70
Notes: Wrong SD. Too narrow. Right SD 5. Wrong SD 2.5.. False rejects.

## Slide 22: Assigned vs. Observed – What is the difference?
Mean assign SD assign
Mean observed SD observed
May or may not be based on fact
Values applied to the L-J Chart
Sources often chosen other than currently observed performance include:
  Past data points before a change
  Cumulative values
  Package insert values
Mathematically calculated using data points (facts)
Answers WHERE WE ARE in terms of accuracy and precision of the method
Specific to your method’s performance

## Slide 23: When Do You Create a New LJ Chart?
Lot changes
  Controls (parallel testing)
  Reagents
  Reference solutions
Major component changed on the instrument
Change in the testing process, usually manufacturer-directed.

## Slide 24: Thank you